## Slide 1: Safe Use and Care of Centrifuges & Rotors
Anthony Curran
Beckman Coulter

## Slide 2: Presentation Objectives
Help you to use your centrifuge systems safely
Ensure maximum life of your centrifuges and accessories
Avoid unbudgeted lab expense

## Slide 3: Beckman: World Leader in Centrifugation Since 1947
The First Ultracentrifuge
Spinco Model L

## Slide 4: Centrifuge Classification

## Slide 5: SAFETY INFORMATION

## Slide 6: Centrifuge accidents produce tremendous amounts of energy
Centrifuges are designed to meet rigorous safety standards. Proper use of rotors and labware, as well as following recommended care procedures found in the Operator Manual, will ensure years of safe operation while avoiding costly accidents

## Slide 7: …hopefully, the last time we’ll see this!

## Slide 8: Approved Systems
Centrifuges and rotors are tested as a System (IEC-1010-2-020) – CE Marking IEC – International Electrical Equipment safety regulations
Only approved Rotor-Centrifuge Systems are proven to be safe
Using non-approved components (including accessories) can result in:
  Centrifuge and/or rotor damage
  Unbudgeted expense
  Personal injury

## Slide 9: Safety Features
Imbalance Detectors
Magnets
Overspeed Disk System
DRIC (Dynamic Rotor Inertia Check)

## Slide 10: Magnets
Used in some of the General Purpose range, all of the High Performance range and all Benchtop Ultracentrifuges.
A rotor identification system prevents the installed rotor from running above its maximum rated speed.
During acceleration the microprocessor checks the magnetic rotor identification (each rotor has a unique angle between the magnets)
If the system identifies a rotor different than the one entered by the user, and the set speed is above the maximum for the identified rotor, the system will reduce the set speed to the maximum for the installed rotor.

## Slide 11: Overspeed Disk
Used on the floor standing ultracentrifuge range

## Slide 12: Overspeed Disk System
This system includes a photoelectric device in the rotor chamber next to the drive spindle and an overspeed disk on the bottom of the rotor
The overspeed disk has alternating light and dark sectors. As the rotor spins, the passage of reflecting and non-reflecting sectors over the photoelectric device generates a series of pulses detected by the electronic circuitry and software.
Each rotor type has a unique number of sectors
After the rotor reaches 600 rpm, the set speed is checked against the overspeed disk

## Slide 13: Dynamic Rotor Inertia Check
A system used across the High Performance and Ultracentrifuge range
DRIC is used to make sure the amount of kinetic energy created is able to be contained
As the rotor accelerates, rotor inertia is measured and the rotor energy is calculated for the speed set by the user
If the calculated rotor energy is determined to be excessive, the instrument will either de-rate the speed to a rotor group appropriate speed or stop the run.

## Slide 14: Pre-Cooling of Rotors
Cool rotor to desired temperature in the centrifuge
Cold Room – Be careful that rotor cavities do not contain condensation

## Slide 15: Installation of rotors

## Slide 16: Rotor attachment: Avanti J-15R
Turn the locking handle clockwise to attach the rotor to the centrifuge spindle.

## Slide 17
Avanti® Rotor Installation
Drive Spindle Assembly

## Slide 18
Drive Spindle Pegs
Avanti® Rotor Installation

## Slide 20
Pin ends indicate the position of the drive pins directly beneath
Avanti® Rotor Installation

## Slide 21
The rotor pins are positioned in the rotor
bore and are
parallel to the BECKMAN
name engraved close to the center of the rotor body
Avanti® Rotor Installation

## Slide 22
Proper
Seating – rotor drive pin between or to the side of spindle pegs,
NOT on top!
Avanti® Rotor Installation

## Slide 23: Lid attachment: Avanti® Series
For rotors with single knobs:
When the rotor is placed in the centrifuge, tighten the knob clockwise.
The single knob locks the rotor to the spindle and attaches the lid to the rotor.
For rotors with dual knobs:
First, turn the lower knob clockwise to attach the lid to the rotor.
When the rotor is placed in the centrifuge, tighten the centre knob clockwise.

## Slide 24: Lid Attachment: Optima® Series Fixed Angle Rotors
Place lid on rotor
Screw on handle until tight

## Slide 25: Proper Handling of Optima® Series Swinging Bucket Rotors
Improper handling can lead to
  mis-hooked buckets
  disturbed gradients
INCORRECT
CORRECT
Proper handling
- hold rotor with both hands
- no mishooks
- no disturbed samples

## Slide 26: Proper Handling of Optima® Series Swinging Bucket Rotors
Mis-hooked buckets are one of the leading causes of rotor mishaps in Ultracentrifuges
Inspect buckets prior to the run

## Slide 27: Rotor Installation: Optima® Series
Place the rotor straight down onto the spindle
There is no tie-down device- run in a vacuum
Some swinging bucket rotors have pins – ensure these are at a 90-degree angle to the drive hub pins

## Slide 28: Use & Care of Rotors

## Slide 29: Why Rotors Fail
Human Error
Stress
Metal Fatigue
Stress Corrosion

## Slide 30: Why Rotors Fail
No.1 - Human Error!
  Failing to tie-down rotors correctly, mismatched rotor parts, mis-hooking swinging buckets, using inappropriate or failed labware, overload (running at faster speed than recommended) and fatigue (metal) –failing to retire a stressed, out of date rotor!
  Never rush!

## Slide 31: Why Rotors Fail
Stress - The centrifugal force created by high rotational speeds generates the load or stress on the metal of the rotor, which causes it to stretch and change in size

## Slide 32: Why Rotors Fail
Metal Fatigue
Any metal structure, following a certain number of stress cycles, will eventually suffer fatigue
When a rotor is repeatedly run up to operating speed and then decelerated, the cyclic stretching and relaxing of the metal cause changes in its microstructure
Depending on how close the rotor is operated to the elastic limit of the metal, and how many alternating stress cycles it experiences, these small changes will eventually become microscopic cracks
With continued use, the cracks will enlarge, and failure of the rotor will ultimately occur.

## Slide 33: Why Rotors Fail
Corrosion—the attack on metal surfaces by moisture, chemicals, or alkaline solutions such as cesium chloride or other salts—is an aluminum rotor’s worst enemy
The resultant pitting reduces the amount of metal available to bear stress during operation at speed
When this occurs in a highly stressed area, such as the bottom of a cell hole in a fixed angle rotor the increased load on the remaining metal results in a so-called stressed concentrator
Repeated use will cause cracks to develop, and failure of the rotor will result
This combination of stress and corrosion is called stress corrosion
Stress corrosion is the primary cause of failure of aluminum rotors, especially as they age.

## Slide 34: Stress corrosion can destroy a rotor
Over time, the combination of tensile load and environment creates stress corrosion in aluminum rotors, significantly reducing the service life rotor
Examples of stress corrosion
Examples of potential results of stress corrosion
Note the anodize protecting the aluminum alloy is completely missing, allowing corrosion to form.

## Slide 35: Points to Remember
Anodized aluminum is corrosion resistant, not corrosion proof
Although titanium and carbon composite rotors are highly corrosion resistant, these rotors will typically include anodized aluminum and other components

## Slide 36: Inspect Regularly
Inspect rotor body and tube cavities
  Pitting, rough spots, cracks or any obvious damage
  Discoloration
  Presence of water
Inspect O-rings for cuts, abrasions or perishing and replace if damaged
Inspect the overspeed disk and replace if damaged
Lubricate rotor and components as indicated in rotor manual
Contact Beckman Coulter Field Service if you have questions
Consult your Operators Manual

## Slide 37: Cleaning procedure
Wash rotor frequently
Remove O-rings before washing
Use mild detergent, such as Solution 555, and soft brushes (both are available in the BCI Rotor Cleaning Kit)
Thoroughly rinse with distilled water
Air-dry upside down - Swinging bucket rotors should be stored with bucket caps removed.
Re-lubricate O-rings with vacuum grease
Re-lubricate metal threads with Spinkote
Consult your Operators Manual

## Slide 38: Proper Cleaning Utensils
BCI Rotor Cleaning Kit

## Slide 39: Correct Brushes
Improper bushes (e.g. brushes with a metal core) can scratch and damage rotor tube cavities.
The brush on the right has a plastic core.

## Slide 40: Correct Brushes

## Slide 41: Respect for Others

## Slide 42: Spindle care

## Slide 43: Decontamination & Sterilization
Radioactive material
  Use a cleaning agent that will not harm anodized aluminum
Biohazardous or Pathogenic contamination
  Follow appropriate procedures outlined by your Laboratory Safety Officer
  Metal rotors and their constituent components can be autoclaved for up to 1 hour at 121◦C
  Cold methods – 70% ethanol and 6% hydrogen peroxide – can be used. Bleach (sodium hypochlorite) may be used but may discolour anodized surfaces.

## Slide 44: O-Rings

## Slide 45: O-Rings
Reasons for O-rings
  Provide frictional resistance to prevent components from separating
  Act as liquid tight seal to prevent leaked contents escaping from rotor
  Provide vacuum tight seal-ensures atmospheric pressure within rotor and tube cavities

## Slide 46: Only Effective Though If Present
Need to be in good condition, clean and greased
Cleaning and Greasing
  Remove with a non-metallic tool
  Wash with mild detergent
  Dry and apply thin layer vacuum grease
Don’t forget to protect metal threads with Spinkote

## Slide 47: O Ring Failure

## Slide 48: Loading
Having assured yourself you have a clean, dry and serviceable rotor – consider carefully what you’re going to put into it!

## Slide 49: Labware Selection & Sample Loading

## Slide 50: Tubes, Bottles and Adapters
A vast selection of tubes, bottles, adapters and other accessories are available for use in centrifuges
Consult the rotor manual for correct use of tubes, bottles, and adapters
Improper labware use is a leading cause of rotor malfunction

## Slide 51: Tubes and Samples
Do your tubes fit?
  do you need caps / sealing mechanism?
    are they assembled / sealed correctly?
  do you need adapters?
    are they loaded correctly?
Do you need to de-rate the speed?
  because of tube material?
  sample volume?
  density of sample/gradient media?
  adapters?
  run temperature?

## Slide 52: Correctly Balance Rotor Loads
At 1,000,000 x g,
1 gram weighs 1,000 kg
The Weight of a Mini Cooper
Which Equals . . .

## Slide 53: For High Performance and General Purpose Benchtop Centrifuges…
Balance opposing sample loads to within 1g for all rotors

## Slide 54: For Ultracentrifuges...
Balance opposing sample loads to within 0.01g for all rotors

## Slide 55: Unusual balancing methods

## Slide 56: Maximum sample density for most rotors is...
Consult your Operator Manual for information on precipitating gradients. Overspeeding self-forming gradients can cause salt to come out of solution with a high density pellet that will settle at the highest force portion of the rotor.
1.2 g/mL for most rotors
1.7 g/mL for VTi and NVT rotors

## Slide 57: Correct Rotor Use depends on Several Factors
Use only as part of approved systems
Correct rotor assembly & installation
Routine inspection & maintenance
Proper use of tubes, adapters, accessories, and tools
Properly balanced loads
Consult your Operators Manual

## Slide 58: Questions ?